## Slide 1
SANS Data Reduction Workflow
Reduction
Parameters
User-Specified Mask Y/N
Bad Tube Mask Y/N
Binning Details
Absolute Scale Factor Y/N
Sample Files
Sensitivity Y/N
Background Y/N
Sample Transmission File or Value
Background Transmission File or Value
Sample Geometry
Beam Center File or Value
Background Geometry
Metadata Output Options Y/N
Sort Variables for Time Slicing
ROIs Y/N
Dark Current or Blocked Beam Y/N
Data Output Options Y/N
Reduce Data
Transmission File or Value?
Sample Transmission Y/N
Yes
No
Sample Transmission File or Value
Calculate Transmission
TS(λ),dTS(λ)
TB(λ),dTB(λ)
Sample Geometry
Set Transmission to Unity
Background Y/N
Background File?
Yes
Background Transmission File or Value
Calculate Transmission
Background Geometry
Background Transmission Y/N
Transmission File or Value?
No
Set Transmission to Unity
Yes
Reduction
Log
Reduction
Log
Reduction
Parameters
Beam Center File or Value
Beam Center Pixel Space
Center Pixel (X,Y)
Reduction
Log
Background FIle
Background Sort Variables for Time Slicing
Read File
Prepare Data
B(X,Y, λ), dB(X,Y, λ)
Reduction
Parameters
Center Pixel (X,Y)
Reduction
Log
1
3
2
Reduction
Parameters
3a
Sample Files
Sort Variables for Time Slicing
Read File
Prepare Data
Center Pixel (X,Y)
D(X,Y,Z,λ), dD(X,Y,Z,λ)
Reduction
Log
Dark Current or Blocked Beam?
Yes
Dark Current or Blocked Beam
Read File
Prepare BB Data OR DC
BB_DC(X,Y), dBB_DC(X,Y)
Reduction
Parameters
Apply Transmission Correction
B1(X,Y,Z,λ), dB1(X,Y,Z,λ)
Normalize for Illuminated Volume
Background Geometry
Reduction
Parameters
B2(X,Y,Z,λ), dB2(X,Y,Z,λ)
3b
Apply Transmission Correction
Normalize for Illuminated Volume
Sample Geometry
D1(X,Y,Z,λ), dD1(X,Y,Z,λ)
Reduction
Parameters
D1(X,Y,Z,λ), dD1(X,Y,Z,λ)
Perform Background Subtractions
EQ-SANS Wavelength Normalization
Absolute (Final) Normalization
4
D2(X,Y,Z,λ), dD2(X,Y,Z,λ)
D4(X,Y,Z,λ), dD4(X,Y,Z,λ)
EQ-SANS?
Yes
F(X,Y,Z,λ), dF(X,Y,Z,λ)
No
F(X,Y,Z,λ), dF(X,Y,Z,λ)
Reduction
Parameters
Absolute Scale Factor
D5(X,Y,Z,λ), dD5(X,Y,Z,λ)
User-Specified Mask Y/N
Instrument
Reduction
Log
User Mask?
Yes
No
Read Mask
Apply Mask
Reduction
Parameters
User-Specified Mask
Mu(X,Y)
5
Reduction
Parameters
D6(X,Y,Z,λ), dD6(X,Y,Z,λ)
D6(X,Y,Z,λ), dD6(X,Y,Z,λ)
R(X,Y,Z,λ), dR(X,Y,Z,λ)
F(X,Y,Z,λ), dF(X,Y,Z,λ)
ROIs Y/N
ROIs?
Yes
Reduction
Parameters
ROIs
Read ROIs
ROIs
Perform ROI Calculations
ROI Results
Reduction
Log
6a
Output ROI Results
Binning Details
Reduction
Parameters
Bin into 2D Q-Space
Bin into 1D Q-Space
F(Q), dF(Q),
σ(Q)
F(QX,QY), dF(QX,QY),
σ(QX,QY)
6
Data Output Options Y/N
Metadata Output Options Y/N
Output 2D Reduced Data
Output 1D Reduced Data
Reduction
Parameters
Special Output Files?
Yes
Output Additional, Special Reduced Data Files
7
Output Metadata?
Yes
Reduction
Log
Output Reduction Log
Metadata Output Options
Reduction
Log
Output Specified Metadata
Bad Tube Mask
User-Specified Mask
Absolute Scale Factor
Background File
Sensitivity File
Sample Transmission Y/N
Background Transmission Y/N
ROIs
Dark Current or Blocked Beam Y/N
Dark Current or Blocked Beam
Metadata Output Options
Data Output Options
Data Output Options
Reduction
Parameters
Transmission Reference File
Sample Transmission File
Background Transmission File
Polarization?
D3(X,Y,Z,λ), dD3(X,Y,Z,λ)
No
Reduction
Log
Full or Half Polarization
Yes
Polarization Type
Full
Half
D3(X,Y,Z,λ), dD3(X,Y,Z,λ)
Apply Full Polarization Corrections from Polarization Efficiencies Calculation
D3(X,Y,Z,λ), dD3(X,Y,Z,λ)
Apply Half Polarization Corrections from Polarization Efficiencies Calculation
D3(X,Y,Z,λ), dD3(X,Y,Z,λ)
Wavelength Cutoffs
TOF Cutoffs
D6(X,Y,Z,λ), dD6(X,Y,Z,λ)
SDD
SDD < 1.5 and Detector Offset > 20 (HFIR)?
Yes
No
Reduction
Log
Create Mask or ROI
Back Plane Tube Numbers
MBT(X,Y)
Apply Mask
R(X,Y,Z,λ), dR(X,Y,Z,λ)
R(X,Y,Z,λ), dR(X,Y,Z,λ)
Detector Offset
TS(λ)=1, dTS(λ)=0
TB(λ)=1, dTB(λ)=0

## Slide 2
SANS Data Reduction Workflow
Set Vectors for Hardware that Defines Instrument Geometry
Reduction
Parameters
Reduction
Log
Reduction
Parameters
Reduction
Log
Source Aperture-to-Sample
Sample Position
Pixel (X,Y,Z,dX,dY) Values
Timing Source-to-Sample
Sample-Detector Distance
User Defined Offsets
Pixel (X,Y,Z,dX,dY)
Sample Position Defined (0,0,0)
Source Aperture (X,Y,Z)
Timing Source (X,Y,Z)
Calculate Zero-Angle Transmission
Reduction
Parameters
Transmission File or Value
File or Value?
File
Transmission Reference File
Value
T0(λ), dT0(λ)=0
Sample Transmission File
Center Pixel (X,Y)
Read File
Prepare Data
Read File
Prepare Data
Integration Method
Reduction
Log
Entered Radius or Area From Beam
Entered Radius
Area from Beam
Integrate Counts in Area
Integrate Counts in Area
Integration Radius
T0,ref(λ), dT0,ref(λ)
T0,sam(λ), dT0,sam(λ)
Divide Spectra
T0,raw(λ), dT0,raw(λ)
Transmission Value
Calculate Transmission Integration Area
ROI(X,Y)
ROI(X,Y)
Create Mask or ROI
Create Mask or ROI
Linear Fit
T0,fit(λ), dT0,fit(λ)
Reduction
Log
Reduction
Log
Calculate Transmission Integration Area
Δλ
SSD
SDD
Source Aperture
Sample Aperture
λ
ROI(X,Y)
Center Pixel (X.Y)
Beam center
Beam spot size in x direction
Beam spot size in y direction
Bottom:
Beam spot size in y direction
top:
Hammouda B and Mildner D.F.R. JAC, 2007, 40, 250-259
See the following reference for the meanings of the symbols
Define the radius of the integration area as 1 pixel smaller than the value of Xmin
The beam is oval not elliptical
Create Mask or ROI
Reduction
Log
Integration Radius
Perform Background Subtraction
BB_DC(X,Y), dBB_DC(X,Y)
B(X,Y, λ), dB(X,Y, λ)
D(X,Y,λ), dD(X,Y, λ)
Background Y/N
Background File?
Dark Current or Blocked Beam?
Dark Current or Blocked Beam Y/N
Yes
Reduction
Parameters
Yes
D1(X,Y,λ), dD1(X,Y, λ)
Mt(X,Y, λ), dMt(X,Y, λ)
Empty Cell Background Y/N
Empty Cell File?
Yes
Background subtraction can be nested multiple layers deep. In its most detailed form commonly used, the subtraction takes the form of the following.
D1(X,Y,λ) = (D(X,Y,λ)-BB_DC(X,Y,λ)) – (Mt(X,Y,λ)-BB_DC(X,Y,λ)) – (B(X,Y,λ)-BB_DC(X,Y,λ)) – (Mt(X,Y,λ)-BB_DC(X,Y,λ))
D(X,Y,λ) is the sample in its cell
B(X,Y,λ) is the matrix that the scattering material of interest is suspended within, such as a solvent background. This may or may not exist. The empty cell may also be used here.
Mt(X,Y,λ) is the empty cell or an empty beam. This may or may not exist. It may also not be appropriate to explicitly subtract the contribution of the empty cell, or it may already be accounted for in B(X,Y,λ).
BB_DC(X,Y) is the blocked beam or dark current. It is always assumed to exist, but it may be negligible.
Prepare Data or DC
Dark Current or Blocked Beam?
Blocked Beam
Dark Current
Read File
Prepare Data
BB(X,Y,λ)
Reduction
Parameters
Dark Current or Blocked Beam
Reduction
Parameters
Blocked Beam
Read File
Normalize to Time
Reduction
Parameters
Dark Current
DC1(X,Y,t)
Flatten TOF
DC2(X,Y)
Use Average Dark Current
Reduction
Parameters
DC3(X,Y)
Average Dark Current?
No
Yes
A
DC4(X,Y)
Calculate Average
Normalize to Average
DC4(X,Y)
DC5(X,Y)
B
Note: Only one of either A or B are output, not both.
Reduction
Parameters
Prepare Data
Read Mask
Set Instrument Geometry
Apply Mask
D(X,Y,t)
D2(X,Y,t), dD2(X,Y,t)
MT(X,Y)
Bad Tube Mask
D1(X,Y,t)
Initial Uncertainty Estimate
From “Read File”
Reduction
Parameters
Reduction
Log
D3(X,Y,Z,t), dD3(X,Y,Z,t)
Correct TOF and Bin into λ
D7(X,Y,Z,λ), dD7(X,Y,Z,λ)
Sensitivity File
S(x,y), dS(x,y)
Read Sensitivity
Correct Sensitivity
Time or Monitor Value
D8(X,Y,Z,λ), dD8(X,Y,Z,λ)
Normalize for Time, Monitor or PC
D9(X,Y,Z,λ), dD9(X,Y,Z,λ)
Solid Angle Correction
D4(X,Y,Z,t), dD4(X,Y,Z,t)
EQ-SANS?
Yes
No
Reduction
Log
Instrument
D5(X,Y,Z,λ), dD5(X,Y,Z,λ)
Reduction
Parameters
Flux Spectrum File
Phi(λ)
Read Flux File
Divide by Flux
D5(X,Y,Z,λ), dD5(X,Y,Z,λ)
D6(X,Y,Z,λ), dD6(X,Y,Z,λ)
Wavelength Cutoffs
TOF Cutoffs
Initial Estimate of Uncertainty
D (X,Y,t)
The uncertainty in D(X,Y,t) is calculated using the following equation.
The equation addresses the question of how to deal with a measured intensity of 0 counts, which is defined to have an uncertainty of 1
D2(X,Y,t), dD2(X,Y,t)
Apply Transmission Correction
T(λ),dT(λ)
Reduction
Log
D(X,Y,λ), dD(X,Y, λ)
Angle-Dependent Transmission Correction?
Angle Dependent Transmission Correction Y/N
No
Yes
Pixel (X,Y,Z,dX,dY)
Sample Position Defined (0,0,0)
Calculate Angles
2θ(X,Y)
T(λ),dT(λ)
,
Reduction
Parameters
T(X,Y,λ),dT(X,Y,λ)
Divide Data
Divide Data
D1(X,Y,λ), dD1(X,Y, λ)
D(X,Y,λ), dD(X,Y,λ)
D1(X,Y,λ), dD1(X,Y,λ)
A
B
Note: Only one of either A or B are output, not both.

## Slide 3
SANS Data Reduction Workflow
Beam Center Pixel Space
M(X,Y)
Reduction
Parameters
D(X,Y,t)
Beam Center Method
Reduction
Log
Bad Pixel Mask File
Direct Beam Data File
Read Mask
Apply Mask
User-Supplied Beam Trap Mask BT(X,Y)
Center Pixel (X,Y)
Read Data
Initial Uncertainty Estimate
D1(X,Y,t)
Flatten TOF
D1(X,Y,t)
dD1(X,Y,t)
D2(X,Y)
dD2(X,Y)
Data Type
Reduction
Parameters
Transmission
Scattering
Direct Beam Data File
Apply Mask
D3(X,Y)
dD3(X,Y)
D3(X,Y)
dD3(X,Y)
Reduction
Parameters
Beam Center Method
Center of Mass
Fit Function
D3(X,Y)
dD3(X,Y)
Fitting Function
Fit Data
Reduction
Parameters
Center Pixel (X,Y)
In Tolerance?
No
Yes
Initialize Fitting Function
In Tolerance?
No
Yes
Center Pixel (X,Y)
Set Center and Smaller Radius for Integration
Read Mask or ROI
Reduction
Parameters
File Name
Read 2D File
Mask or ROI?
Mask or ROI?
Mask
ROI
Is Mask
Is ROI
F(X,Y)
Append Metadata
M(X,Y)
Append Metadata
R(X,Y)
A
B
Note: Only one of either A or B are output, not both.
Read Raw Data
Reduction
Parameters
File Name
Nexus or XML?
XML
Nexus
Read Nexus
Read XML
Metadata
Neutron Data
Sort Metadata
IPTS Information
Instrument Configuration
Sample Environment
Tag Metadata
Tag Metadata
Tag Metadata
Reduction
Log
D(X,Y,t)
Create Workspace
Write Reduced Data
Reduction
Parameters
User-Specified File Name
1D or 2D?
1D
2D
Reduction
Log
Reduced Data
1D or 2D?
Write CANSAS1D
Write 1D ASCII
Write Nexus
Write 2D ASCII
Files used in Reduction
Process Start Time
Process End Time
Reduction Logfile Name and Path
Generate Header Strings
Bin into 2D Q-space
Reduction
Parameters
Reduction
Log
Binning Details
F(X,Y,Z,λ),
dF(X,Y,Z,λ)
Number of QX=NX
Number of QY =NY
Pixel (X,Y,Z,dX,dY)
Sample Position Defined (0,0,0)
Calculate Qmax of Corners
λmin
QX,min
QY,max
QY,min
QX,max
Empty (NX,NY) Data Arrays: N, Q, I, dI, and σQ
Reduction
Log
EQ-SANS?
Instrument
No
Yes
Details regarding the binning into 2D Q-space are contained in the document heller_qbinning_030518.docx.
The information for calculating the uncertainty in Q, σQ, is contained in the document SANS Instrument Resolution_LDS_inprogress.docx.
ΔλVS
SSD
SDD
Source Aperture
Sample Aperture
λ
RMS(λi).. RMS(λN)
SSD
SDD
Source Aperture
Sample Aperture
λi.. λN
Δλi.. ΔλN
F(QX,QY), dF(QX,QY),
σ(QX,QY)
F(QX,QY), dF(QX,QY),
σ(QX,QY)
Reduction
Parameters
Reduction
Log
Binning Details
Extract Qmax
Details regarding the binning into 1D Q-space are contained in the document heller_qbinning_030518.docx.
During discussions with the team, it was suggested that the various kinds of line, box and wedge averages could be thought of as a set of ROIs.
Number of Q = N
1D Binning Style
Determine Q-values
Empty (N) Data Arrays: N, Q, I, dI, and σQ
Fill Q-array
Empty (N) Data Arrays: N, dI, and σQ
Q-value Array
Wedge, Line, Box ROIs
R1(X,Y)..RN(X,Y)
F(Q), dF(Q),
σ(Q)
Bin into 1D Q-space

## Slide 4
SANS Data Reduction Workflow...operations outside the normal workflow
Barscan Calibration and Detector Map
Bar@Z0 File
Bar@Z0 Z0(X,Y,t)
Bad Pixel Mask File
Read Mask
Read File
Apply Mask
Mask(X,Y)
Z0’(X,Y)
Flatten TOF
Bar@Z0 Z0(X,Y)
Bar@ZN File
Bar@ZN ZN(X,Y,t)
Read File
Flatten TOF
Bar@ZN ZN(X,Y)
Apply Mask
ZN’(X,Y)
Sort into Tubes 0..M
Sort into Tubes 0..M
Bad Pixel Mask File
Read Mask
Mask M(X,Y)
Flood Field File
F(X,Y,t)
Read File
Apply Mask
F’(X,Y)
Flatten TOF
F(X,Y)
Sort into Tubes 0..M
F’M(Y)
F’0(Y)
Calculate Counts / Unit Length
C0
CM
Fit to Determine Δwidth
Calculate M Tube Widths
TM,N(Y)
T0,0(Y)
Calculate Average in Tube
Ave0,0
AveM,N
Find Lower Edge
Convert to Physical Coordinate
BtmM(N)
Btm0(N)
Fit Fifth Order Polynomial
Calculate M Tube Horizontal Positions
H0
HM
A0,0
A0,5
AM,0
AM,5
W0
WM
V0,0
V0,K
VM,0
VM,K
Calculate K Pixel Vertical Positions
Calculate K Pixel Vertical Widths
PV0,0
PV0,K
PVM,0
PVM,K
Aggregate M,K map of positions and widths
Map(X,Y,dX,dY)
Write Map File
Set Grid Values in Rectangle to 0
Template Data T(X,Y)
Mask or ROI?
Draw Shape
Is the region specification complete?
Create a Mask or ROI
Upper Left Corner
Template Grid G(X,Y)
Lower Right Corner
Tube Number
Pixel ID
Grid Gi(X,Y)
Apply a Specified Rectangle
Set Grid Values in Tube to 0
Grid Gi(X,Y)
Apply a Specified Tube
Set Grid Value in Pixel to 0
Grid Gi(X,Y)
Apply a Specified Pixel
Set Grid Values in Shape to 0
Grid Gi(X,Y)
Apply a Drawn Shape
Beam Center X,Y
Beamtrap Radius
Set Grid Values within Radius of Beam Center to 0
Apply the Beamtrap
Grid Gi(X,Y)
Yes
No
Mask or ROI?
Mask
ROI
Grid Gi+1(X,Y)
Grid Gfinal(X,Y)
Write Mask or ROI
Mask M(X,Y)
ROI R(X,Y)
Invert Values in Grid
Integration Radius
Set Grid Values within Defined Radius to 0
Apply the Transmission Integration Area
Grid Gi(X,Y)
Calculate Transmission Integration Area
Prepare a Detector Sensitivity File
Flood Field File(s)
Flood Field I(X,Y,t),dI(X,Y,t)
Bad Pixel Mask File
Read Mask
Read File
Solid Angle Correction
Apply Mask
Mask M(X,Y)
I1(X,Y,t), dI1(X,Y,t)
I2(X,Y,t), dI2(X,Y,t)
Flood Field Transmission File
Calculate Transmission
TFF(λ),dTFF(λ)
Apply Transmission Correction
I3(X,Y,t), dI3(X,Y,t)
Calculate Average
Iave, dIave
Apply Thresholds
Draw Patch
Fit over patch and apply values
I4(X,Y,t), dI4(X,Y,t)
I5,patch(X,Y,t), dI5,patch(X,Y,t)
Minimum Threshold
Maximum Threshold
Single sensitivity file
Apply Beamstop Mask
Average Data Sets
I5,set(X,Y,t), dI5,set(X,Y,t)
I6(X,Y,t), dI6(X,Y,t)
Multiple sensitivity files
Normalize to 1.0
S(X,Y,t), dS(X,Y,t)
Write Sensitivity File
Calculate Average
Iave, dIave
Single or Multiple Sensitivity Files
Flood Field File(s)
Prepare Data
Merge 1D Data
FIle 1
File N
Read 1D Data File
DN(Q,I,dI,dQ)
D1(Q,I,dI,dQ)
Reference File No. M
Qoverlap1,2,min
QoverlapN-1,N,min
Integrate from Qoverlap,min to Qoverlap,max
F0
FN
Scale Relative to Mth value
F0/FM
FN/FM
Scale Intensity by Factor
DN(Q,(F0/FM)*I,(F0/FM)*dI,dQ)
D1(Q,(F0/FM)*I,(F0/FM)*dI,dQ)
Aggregate into Single 1D Profile Sorted by Ascending Q
Dmerge(Q,I,dI,dQ)
Output 1D Reduced Data
Determine Polarization Correction Factors
Full or Half Polarization?
Full or Half Polarization?
Half
Full
f
Pm
Note: These are either user-supplied or calculated prior to being here
Lisa and Mike owe us Math…
S+
S-
Note: These are either user-supplied or calculated prior to being here
A
B
Note: Only one of either A or B are output, not both.
…or Math and Pizza
T
Pm
f
Pa
S+,+
S+,-
S-,-
S-,+
Qoverlap1,2,max
QoverlapN-1,N,max

## Slide 5
SANS Data Reduction Workflow...operations outside the normal workflow
Determine Illuminated Volume
Project Beam Through Sample
Calculate or Manual Input?
Illuminated Volume Method
Illuminated Volume
Reduction
Parameters
Reduction
Parameters
Sample Volume S(X,Y,Z)
Reduction
Log
Source Aperture (X,Y,Z)
Sample Aperture
Illuminated Volume
Illuminated Volume
Reduction
Log